sell Series <Version 5.9, PCB no.2L100I>

##### Notice: Before entering the setting procedures, please disassemble the housings of the scale.  Then find the JP1 jumper (located in the above of U3 IC93C66 on Main PCB, and adjust it in order to enable the setting mode.

( Modify Specification (

| Instructions | Display | Descriptions |
| --- | --- | --- |
| Press [RESTO] key and switch on the scale at the same time. Do not release your hand from [RESTO] key till Display shows version of scale. | WEIGHT Display: 5.9 UNIT PRICE Display: - 1- | Version of scale |
| Press [TARA] key and hold till WEIGHT Display shows SEtUP. | WEIGHT Display: SEtUP | Entering SEtUP mode |
| 3. Press [2] key once, and then the WEIGHT Display will show SPEC. | WEIGHT Display: SPEC. | Enter the setting mode of SPEC. (Modify Specification). |
| 4. Mode 1: Setup Max Capacity under mode 1. |
| 4-1. Press [C] key to select mode 1. | WEIGHT Display: 1. | Select mode 1. |
| 4-2. Press [ . ] key to enter. | WEIGHT Display: XXXXXX | Enter mode 1. |
| 4-3. Press [ . ] key to shift the flashing digit & [C] key to setup the flashing digit. e.g. For 6/15kg x 2/5g, the max capacity would be 15 kg. | WEIGHT Display: XXXXXX ↓ 00015 | 1. Press [ . ] key to shift the flashing digit. Stop at the digit you needed to correct figure. 2. Press [C] key to select the figure from 0~9. Correct the figure to the value of max capacity. |
| 4-4. Press [RESTO] key to confirm. | WEIGHT Display: good ↓ 1. | WEIGHT Display will flash good twice and the indicator / scale will go back to mode 1. automatically. |
| 5. Mode 2: Setup Max Display Resolution under mode 2. |
| 5-1. Press [C] key to select mode 2. | WEIGHT Display: 2. | Select mode 2. |
| 5-2. Press [ . ] key to enter. | WEIGHT Display: XXXXXX | Enter mode 2. |
| 5-3. Press [ . ] key to shift the flashing digit & [C] key to setup the flashing digit. e.g. For 6/15kg x 2/5g, the max display resolution would be 7500. | WEIGHT Display: XXXXXX ↓ 007500 | 1. Press [ . ] key to shift the flashing digit. Stop at the digit you needed to correct figure. 2. Press [C] key to select the figure from 0~9. Correct the figure to the value of max display resolution. |
| 5-4 Press [RESTO] key to confirm. | WEIGHT Display: good ↓ 2. | WEIGHT Display will flash good and the indicator / scale will go back to mode 2. automatically |
| 6. Mode 3: Correct sensitivity (minimum capacity) & decimal point under mode 3. |
| 6-1. Press [C] key to select mode 3. | WEIGHT Display: 3. | Select mode 3. |
| 6-2. Press [ . ] key to enter. | WEIGHT Display: XXXXXX | Enter mode 3. |
| 6-3-1. Press [ . ] key to shift the flashing digit. 6-3-2. Press [C] key to setup the flashing digit. 6-3-3. Press [ZERO] key to put or withdraw the decimal point. e.g. For 6/15kg x 2/5g, the sensitivity would be 0.002kg (2g). | WEIGHT Display: XXXXXX ↓ 000.002 | 1. Press [ . ] key to shift the flashing digit. Stop at the digit you needed to correct the figure. 2. Press [C] key to select the figure from 0~9. Correct the figure to the value of sensitivity. 3. Press [ZERO] key to put or withdraw the decimal point at the flashing digit as XXX.XX (based on kg mode). |
| 6-4. Press [RESTO] key to confirm. | WEIGHT Display: good ↓ 3. | WEIGHT Display will flash good and the indicator / scale will go back to mode 3. automatically. |
| 7. Mode 4: Set up the Decimal Point on Weight Display under mode 4. |
| 7-1. Press [C] key to select mode 4. | WEIGHT Display: 4. | Select mode 4. |
| 7-2. Press [ . ] key to enter. | WEIGHT Display: XXX.XXX | Enter mode 4. |
| 7-3-1. Press [ . ] key to shift the flashing digit. 7-3-2. Press [ZERO] key to put or withdraw the decimal point. | WEIGHT Display: 000.000 ↓ 000.000(NO RESTO) | 1. Press [ . ] key to shift the flashing digit. Stop at the digit you needed to correct the figure. 2. Press [ZERO] key to put or withdraw the decimal point at the flashing digit. |
| 7-4. Press [RESTO] key to confirm. | WEIGHT Display: good ↓ 4. | WEIGHT Display will flash good and the scale / indicator will go back to mode 4. automatically. |
| 8. Press [RESTO] key twice to finish the setting, the indicator / indicator will restart automatically with new setting. |
| 9. Please refer to Calibration Procedure as below and calibrate the indicator / scale again before using. |

Series <Version 5.9, PCB no.2L100I>

##### Notice: Before entering the setting procedures, please disassemble the housings of the scale.  Then find the JP1 jumper (located in the above of U3 IC93C66 on Main PCB, and adjust it in order to enable the setting mode.

( Weight Calibration (

| Instructions | Display | Descriptions |
| --- | --- | --- |
| 1. Press [RESTO] key and switch on the scale at the same time. Do not release your hand from [RESTO] key till Display shows version of scale. | WEIGHT Display: 5.9 UNIT PRICE Display: - 1- | Version of scale. |
| 2. Press [TARA] key, hold till WEIGHT Display shows SEtUP | WEIGHT Display: SEtUP | Entering SEtUP mode |
| 4. Press [TARA] key once, WEIGHT Display shows 000.000 | WEIGHT Display: 000.000 | Enter calibration mode. |
| 5. Press [ . ] key to shift the flashing digit. | WEIGHT Display: 000.000 ↓ 000.000 | The flashing digit will be shifted once you press [ . ] key. Stop at the digit you need to correct the figure. |
| 6. Press [C] key to set the value of standard weight. | WEIGHT Display: 000.000 ↓ 010.000 | ** Example: If you use 10kg standard weight to calibrate the scale, press [ . ] key to shift the flashing digit to 000.000 and press [C] key to set the figure to 010.000 |
| 7. Press [RESTO] key to confirm. | WEIGHT Display: 6XXXXX | WEIGHT Display will identify internal Off Set value. |
| 8. Put standard weight on the platter/platform. And then Weight Display will show internal Span value. | WEIGHT Display: XXXXXX | Identify internal Span value. |
| 9. Wait for a moment till Span value is stable (2~3 unstable counts are acceptable. Do not wait too long), and then press [RESTO] key once to confirm. Weight Display will flash “good” twice. | WEIGHT Display: XXXXXX ↓ good | Identify internal Span value is stable, and save the data into IC93C66. !! Notice: When LCDs are counting 11111 ~ 99999, please remove the standard weight from the platter. |
| 10. The scale / indicator will restart automatically. | WEIGHT Display: 000.000 (re-start automatically) | Finish Calibration. |

# Text boxes

# JP1 Setting:

JMP1

DISABLE

JMP1

ENABLE

# JP1 Setting:

JMP1

DISABLE

JMP1

ENABLE